# What is ECG/EMG

By measuring the voltage across the body amplifying it, you can get heart rate or muscle activity. Depending upon the band-pass filter frequency you get heart rate or muscle activity.

## Heart rate: (ECG)

Signals of the heart reach UP TO ±5mV  [1]

0.5Hz to 150Hz

Electrodes must be placed equidistant from the heart for 1 Lead ECG. “Einthoven triangle”

## Muscle activity

Amplitude varies depending upon age, sex, strength, etc. Try to maximize the resolution to the best of our ability as we don’t know.

- The way EMG works is you filter raw signal 20Hz - 500Hz

- Rectify it, remove the negative components

- Put a HP filter to extract the envelope 3.6Hz

## Gain and filter

Ideally, we would be able to tune all this with programmable gain and filter. I don’t know if there’s a programmable filter.

Max30003 does not have programmable filter.

Hardware components

We want the highest resolution possible for the ADC especially since our signal is so small. Delta-sigma ADC

Programmable gain, programmable filter

LDO power supply

# Connectors

## Requirements

We want minimal noise introduced, strong enough connections such that they don’t fall off, yet detachable and replicable.

# Sources

Source 1: “High-resolution portable bluetooth module for ECG and EMG acquisition

”https://www.sciencedirect.com/science/article/pii/S2001037025001394

Source 2: https://learn.sparkfun.com/tutorials/getting-started-with-the-myoware-20-muscle-sensor-ecosystem/myoware-20-muscle-sensor